Rev A

QUICK REFERENCE CARD

Targeted Single-cell DNA + RNA Fusion Sequencing Using the Tapestri® Platform v3

# IMPORTANT

- Ensure instrument is using the current v3 firmware.

- Vortex all reagents unless directed otherwise.

- Thaw -20 °C reagents on ice.

- Avoid sources of static and pipette slowly and carefully when handling emulsions.

- Provide debris-free cell suspension with >90% viability.

- Final cell concentration needs to be ~3,000 cells/μL

- Equilibrate Encapsulation and Barcoding oils for 30 minutes prior to use.

- Always use a PCR skirt.

# Prepare Single Cell Suspension

- Use DPBS to dilute cells to ~ 3,000 cells/μL in a total volume of at least 35 μL.

IMPORTANT   Adding a negative cell line at 5% is highly recommended.

# Encapsulate Cells

- Mount the Base Plate onto the Tapestri Instrument.

- Place 0.2 mL emulsion-safe PCR tube into the middle of the slot at the left side of the Base Plate.

- Place the DNA cartridge onto the base plate.

- Preheat the thermal cycler with the Lysis/Digest protocol (hold at 50C).

- Preheat the RNA Fusion reverse primer pool to 70 °C for 10 minutes, then store on ice until ready to use.

| Step | Temperature | Time |
| --- | --- | --- |
| 1 | 50 °C | 120 min |
| 2 | 80 °C | 10 min |
| 3 | 4 °C | HOLD |

- In a 0.2 mL PCR tube, combine 2 μL RNA Lysis Mix with 18 μL nuclease-free water. Gently mix by pipetting up and down 5 times. Do not vortex. Store on ice.

- In a new tube, prepare RT + Lysis Mix:

| Reagent | Volume (μL) |
| --- | --- |
| RNA Enzyme Buffer | 12 |
| RNA Reagent 1 | 3 |
| RNA Reagent 3 | 3 |
| RNA Reagent 4 | 3 |
| RNA Reagent 2 | 5.4 |
| RNA Enzyme | 3 |
| RNA Fusion reverse primer pool | 6 |
| Nuclease-free water | 23.3 |
| Diluted RNA Lysis Mix | 1.3 |
| Total Volume | 60 |

IMPORTANT   Add the Diluted RNA Lysis Mix immediately before loading the cartridge.

- Pipette 60 μL of RT + Lysis Mix into

reservoir 1.

- Gently pipette mix the Cell Suspension

then pipette 35 uL (3,000 cells/μL) into reservoir 2.

- Pipette 200 μL of Encapsulation Oil into reservoir 3.

- Apply DNA Gasket on top of the cartridge and close the instrument lid.

- Run the Cell Encapsulation program by pressing Step 1: Encapsulation on the Tapestri Instrument touch screen.

- Once program completes, press DONE and remove the cartridge and collection tube from the Base Plate.

- Use a gel loading tip to carefully remove excess oil from the bottom layer of the tube leaving a total of 100 µL of emulsion + oil

(approximately the middle point of the tube as it narrows to the conical base).

IMPORTANT   This is not a stopping point. Immediately move onto RT + Lysis.

# RT + Lysis

- Immediately place the sample tube into the preheated thermocycler and run the Lysis/Digest protocol.

| Step | Temperature | Time |
| --- | --- | --- |
| 1 | 50 °C | 120 min |
| 2 | 80 °C | 10 min |
| 3 | 4 °C | HOLD |

- Once complete, store the lysed and digested sample at 4 °C until required in the next step. Proceed within 2 hours.

# Barcode Cells

PRIME THE DNA CARTRIDGE

- Place eight emulsion-safe PCR tubes

into the eight slots at the bottom of the

Base Plate.

- Mount the DNA Cartridge back onto the Base Plate.

- Pipette 200 μL of Electrode Solution into each reservoir 4 and 500 μL of Electrode Solution into each reservoir 5.

- Apply DNA Gasket on top of the cartridge and close instrument lid.

- Run the Priming program by pressing Step 2: Barcoding on the Tapestri Instrument touch screen.

- In the meantime, retrieve the Barcoding Oil and Barcoding Beads v3 () and leave at room temperature (protect from light) and prepare 300 μL Barcode Mix:

| Reagent | Volume (μL) |
| --- | --- |
| Barcoding MM v3 | 290.4 |
| DNA forward primer pool | 5 |
| DNA reverse primer pool | 3.1 |
| RNA Fusion forward primer pool | 1.5 |
| Total Volume | 300 |

- Briefly vortex the Barcode Mix, quick-spin to collect the contents, and store on ice.

LOAD THE DNA CARTRIDGE

- Quick-spin Barcoding Beads v3 to collect contents. Take 67 μL of the prepared Barcode Mix and add it to the barcoding bead tube.

- Retrieve the lysed and digested sample from the thermocycler.

- Slowly pipette the entire cell encapsulation sample (~100 μL), including any oil at the bottom into reservoir 6.

- Vortex Barcoding Beads v3 at full speed for 1 minute. Pipette 250 μL of Barcoding Beads v3  into reservoir 7.

- Pipette 200 μL of Barcode Mix into reservoir 8.

- Pipette 1.25 mL of Barcoding Oil into reservoir 9.

- Apply DNA Gasket on top of the cartridge and close the instrument lid.

- Run the Cell Barcoding program by pressing NEXT on the Tapestri Instrument touch screen.

- Once the program completes, press DONE and remove the Base Plate together with the cartridge and the eight collection tubes.

- Remove the DNA Cartridge from the Base Plate.

- Use a gel loading tip to carefully remove excess oil from the bottom layer of all eight tubes, leaving a total of 100 µL of emulsion + oil (approximately the middle point of the tube as it narrows to the conical base) per tube.

# Targeted PCR Amplification

- Transfer the sample tubes to the thermocycler, and run the Targeted PCR protocol based on the table located on the next page

| Amplicon Number |  |  | 20-100 | 101-200 | 201-300 | >300 |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Step | Ramp | Temp | Time | Time | Time | Time | Cycle |
| 1 | 4 °C/s | 98 °C | 6 min | 6 min | 6 min | 6 min |  |
| 2 | 1 °C/s | 95 °C | 30 sec | 30 sec | 30 sec | 30 sec | 11 |
| 3 |  | 72 °C | 10 sec | 10 sec | 10 sec | 10 sec |  |
| 4 |  | 61 °C | 3 min | 4.5 min | 6 min | 9 min |  |
| 5 |  | 72 °C | 20 sec | 20 sec | 20 sec | 20 sec |  |
| 6 | 1 °C/s | 95 °C | 30 sec | 30 sec | 30 sec | 30 sec | 13 |
| 7 |  | 72 °C | 10 sec | 10 sec | 10 sec | 10 sec |  |
| 8 |  | 51 °C | 3 min | 4.5 min |  |  |  |
| 9 |  | 72 °C | 20 sec | 20 sec | 20 sec | 20 sec |  |
| 10 | 4 °C/s | 72 °C | 2 min | 2 min | 2 min | 2 min |  |
| 11 |  | 4 °C | HOLD | HOLD | HOLD | HOLD |  |

IMPORTANT   Make sure to set up two separate cycle programs (11 cycles and 13 cycles)  each with a ramp rate of 1 ºC/s.

BREAK EMULSIONS

- Add 15 μL of Extraction Agent to each tube. Briefly vortex and spin for 20 seconds.

- Incubate at room temperature for 3 minutes.

- Add 50 μL of nuclease-free water to each tube.

- Briefly vortex and spin for 10 seconds.

- Pipette the total volume of the aqueous top layer from each tube into two new 1.5 mL DNA LoBind Eppendorf tubes (pool contents from tubes 1-4 in one 1.5 mL tube and tubes 5-8 in a second 1.5 mL tube). Do not transfer any oil or Barcoding Beads.

- Spin down 1.5 mL tubes at 3000 g for 5 minutes. There should be two layers, the clear aqueous layer on top and an opaque layer on the bottom.

- Carefully transfer the clear supernatant from each tube into a new 1.5 mL DNA LoBind Eppendorf tube.

- Bring the total volume in each tube up to 168 uL with nuclease-free water.

- Proceed to next step or store at -20 °C.

# Digest PCR Product

- Add 20 uL of DNA Clean Up Buffer and 12 μL DNA Clean Up Enzyme to each sample tube for a total of 200 uL.

- Briefly vortex and spin down the tube.

- Transfer tube to a thermo mixer or heat block and incubate at 37 °C for 60 minutes.

- Immediately continue to the next step.

# Clean Up PCR Product

NOTE   Equilibrate AMPure XP reagent to room temperature.

Prepare 5 mL fresh 80% ethanol using nuclease-free

water.

### Spin down sample tube for 20 seconds. If a pellet is visible, transfer clear aqueous solution to a new tube, being careful not to disturb the pellet. Add nuclease-free water if any evaporation occurred for a total volume of 200 μL.

- Thoroughly vortex AMPure XP reagent at high speed immediately prior to usage.

- Add 144 μL (0.72X) of AMPure XP reagent to tube. Vortex for 5 seconds and quick-spin to collect contents.

- Incubate tube at room temperature for 5 minutes.

- Place on magnet and wait 5 minutes for the beads to separate from solution.

- Without removing the tube from the magnet, remove the clear liquid from tube and discard.

- Carefully add 800 uL of the freshly prepared 80% ethanol, wait 30 seconds. Without removing the tube from the magnet, remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 53 once, for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from tube without disturbing the beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 5 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet. Add 50 uL of low TE into tube. Vortex and quick-spin to collect the contents.

- Incubate the tube at room temperature for 2 minutes.

- Place the tube onto the magnet and wait for at least 2 minutes or until the solutions are clear.

- Transfer 50 μL of purified PCR product to a new 0.2 mL PCR tube. There will be two 0.2 mL PCR tubes per Tapestri run.

# Separate DNA and RNA Fusion amplicons

## RNA Fusion AMPLICON CLEANUP

PREPARE STREPTAVIDIN BEADS

- Thoroughly vortex Streptavidin beads. Transfer 60 μL of Streptavidin Beads to a new 1.5 mL DNA LoBind Eppendorf tube.

- Place on a magnet and wait 2 minutes for the beads to separate from the solution.

- Remove the supernatant and discard. Wash the beads with 1 mL of nuclease-free water. Wait 1 minute with the tube on the magnet for solution to clear then remove the supernatant.

- Repeat Step 64 once for a total of two washes.

- Resuspend the beads in 30 μL of 1X Wash Buffer (listed below).

| Reagent | Final concentration |
| --- | --- |
| Tris-HCl pH 7.5 | 5 mM |
| EDTA | 0.5 mM |
| NaCl | 1 M |

ISOLATE RNA Fusion AMPLICONS

- Retrieve the two tubes from the Clean Up PCR Product section that contains the DNA and RNA Fusion amplicons.

- To each tube, add 2 μL of TS biotin oligo. Briefly vortex and quick-spin.

- Incubate at 96 °C for 5 minutes.

- Transfer the tubes immediately onto ice and incubate for 5 minutes.

- Add and mix 15 μL of Streptavidin Beads resuspended in 1X Wash Buffer to each Biotin Oligo-treated sample tube.

- Incubate for 30 minutes at room temperature.

- Place on magnet, wait 5 minutes for the beads to separate from the solution.

- Transfer the supernatant to new 0.2 mL PCR tubes. These tubes contain the DNA amplicons.

- Wash the Streptavidin Beads with 33 uL of nuclease-free water and add this supernatant to the 0.2 mL PCR tube containing the DNA amplicons.

- In each tube, resuspend the streptavidin beads in 15 μL of nuclease-free water. The RNA Fusion amplicons are bound to the streptavidin beads. Proceed to Library PCR or store at 4 °C.

DNA AMPLICON CLEANUP

- Take the two 0.2 mL PCR tubes containing the DNA amplicons and bring each up to 100 uL with nuclease-free water if needed.

- Add 72 μL (0.72X) of AMPure XP reagent to the tubes.

- Vortex for 5 seconds and quick-spin to collect the contents.

- Incubate tube at room temperature for 5 minutes.

- Place on magnet, wait 5 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of the freshly prepared 80% ethanol, wait 30 seconds. Without removing the tube from the magnet, remove the ethanol without disturbing the AMPure XP beads

- Repeat Step 82 once, for a total of two washes

- Keeping the tube on the magnet, remove all residual ethanol from tube without disturbing the beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet. Add 15 μL of nuclease-free water. Vortex and quick-spin to collect the contents.

- Incubate the tube at room temperature for 2 minutes.

- Place the tube onto the magnet and wait for at least 2 minutes or until the solutions are clear.

- Transfer 15 μL of purified DNA amplicons to new 0.2 mL PCR tubes.

- Store the purified PCR product on ice and proceed to the next step, or store at -20 °C long term.

# Library PCR

- Set up two different Library PCR reactions, one for the DNA Library and one for the RNA Fusion Library as follows:

IMPORTANT  	Ensure DNA Library Index Primers are used for DNA amplicons, and corresponding TS Library Index Primers are used for RNA Fusion amplicons. Although different library index primers are used, sample indexes should be identical.

|  | DNA | RNA Fusion |
| --- | --- | --- |
| Reagent | Volume [μL] |  |
| Library Mix | 25 | 25 |
| DNA Library Index Primer | 10 | - |
| TS Library i7 Primer | - | 5 |
| TS Library i5 Primer | - | 5 |
| DNA amplicons | 15 | - |
| Resuspended Streptavidin Beads containing RNA Fusion amplicons | - | 15 |
| Total Volume | 50 | 50 |

- Vortex and quick-spin to collect the contents.

- Transfer the sample tubes to a thermocycler and run the following Library PCR protocol:

| Step | Temperature | Time | Cycle |
| --- | --- | --- | --- |
| 1 | 95 °C | 3 min |  |
| 2 | 98 °C | 20 sec | 12 for DNA Library 14 for RNA Fusion Library |
| 3 | 62 °C | 20 sec |  |
| 4 | 72 °C | 45 sec |  |
| 5 | 72 °C | 2 min |  |
| 6 | 4 °C | HOLD |  |

# Clean Up Library PCR Product

NOTE  	Equilibrate AMPure XP reagent to room temperature. Thoroughly vortex AMPure XP reagent for 15 seconds at high speed.

DNA LIBRARY CLEANUP

- Add nuclease-free water to the sample tubes containing the DNA library to bring up to 50 uL if evaporation occurred.

- Add 0.69X of AMPure XP reagent to the 50 μL DNA library tubes. Vortex for 10 seconds and quick-spin to collect the contents.

- Incubate the tube at room temperature for 5 minutes.

- Place on the magnet, wait 5 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of freshly prepared 80% ethanol, wait 30 seconds, and remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 100 once for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from the tube without disturbing the AMPure XP beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet and add 50 μL of nuclease-free water. Vortex and quick-spin to collect the contents.

- Incubate for 2 minutes.

- Place the tubes on the magnet and wait 2 minutes or until the solution is clear.

- Transfer 50 μL of the DNA library from each tube to new 0.2 mL PCR tubes.

- Add 0.69X of AMPure XP reagent to the 50 μL sample tube.

- Incubate the tube at room temperature for 5 minutes.

- Place on the magnet, wait 5 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of freshly prepared 80% ethanol, wait 30 seconds, and remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 111 once for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from the tube without disturbing the AMPure XP beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet and add 12 μL low TE. Vortex and quick-spin to collect the contents.

- Incubate the tube at room temperature for 2 minutes.

- Place the tube onto the magnet and wait for at least 2 minutes or until the solution is clear.

- Transfer 12 μL of DNA library from each tube to new 0.2 mL PCR tubes or 1.5 mL DNA LoBind Eppendorf tube. Store at -20 °C for long-term storage.

RNA Fusion LIBRARY CLEANUP

- Place the tubes containing RNA Fusion Library on the magnet and wait 2 minutes for Streptavidin Beads to separate from the solution.

- Without removing the tube from the magnet, transfer 50 μL of supernatant to new 0.2 mL PCR tubes.

- Add 34.5 uL (0.69X) of AMPure XP reagent to the 50 μL sample tube.

Vortex for 10 seconds and quick-spin to collect the contents.

- Incubate the tube at room temperature for 5 minutes.

- Place on the magnet and wait 5 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of freshly prepared 80% ethanol, wait 30 seconds, and remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 126 once for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from the tube without disturbing the AMPure XP beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet and add 50 μL of nuclease-free water. Vortex and quick-spin to collect the contents.

- Incubate for 2 minutes.

- Place the tubes on the magnet and wait 2 minutes or until the solution is clear.

- Transfer 50 μL of the RNA Fusion library from each tube to new 0.2 mL PCR tubes.

- Add 34.5 uL (0.69X) of AMPure XP reagent to the 50 μL sample tube.

- Incubate the tube at room temperature for 5 minutes.

- Place on the magnet, wait 5 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of freshly prepared 80% ethanol, wait 30 seconds, and remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 138 once for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from the tube without disturbing the AMPure XP beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet and add 12 μL low TE. Vortex and quick-spin to collect the contents.

- Incubate the tube at room temperature for 2 minutes.

- Place the tube onto the magnet and wait for at least 2 minutes or until the solution is clear.

- Transfer 12 μL of RNA Fusion library from each tube to new 0.2 mL PCR tubes or 1.5 mL DNA LoBind Eppendorf tube. Store at -20 °C for long-term storage.

# Quantify and Pool Library

- Run 1 μL of each library on a High-Sensitivity Bioanalyzer chip or equivalent. Repeat AMPure cleanups if necessary. Once the libraries are clean, quantify by qPCR. Normalize libraries to 2 nM for sequencing.

IMPORTANT 	Refer to the Tapestri Single-Cell DNA Sequencing v3 User Guide (MB05-0017) for additional information on how to accurately quantify Tapestri libraries that include large off-target products.

# Sequence Tapestri Single-cell DNA and RNA Fusion Libraries

## IMPORTANT

- The final DNA library consists of target-specific amplicons ranging from 350 - 550 bp with a peak at 460 bp.

- The final RNA Fusion library consists of target-specific amplicons ranging from 350 - 850 bp.

- Refer to the Tapestri Single-Cell DNA Sequencing v3 User Guide (MB05-0017) for sequencing recommendations.

- DNA libraries should be sequenced at 100X per amplicon per cell.

- RNA Fusion libraries should be sequenced at 5000X per amplicon per cell (RNA Fusion panel – [N] amplicons).

For technical support visit www.missionbio.com/support or email support@missionbio.com.

© 2023 Mission Bio, Inc. All rights reserved. Duplication, public display, public performance, modification, or any other use of all or any portion of this document without the express prior written consent of Mission Bio is strictly forbidden. “Mission Bio” and the “MB” logo are trademarks of Mission Bio and may not be used for any purpose without the express prior written consent of Mission Bio. All other trademarks are the property of their respective owners. Nothing in this documentation shall constitute a warranty, express or implied, as to the products (whether a product of Mission Bio, its affiliates, or any third party) or any protocols described herein, whether as to performance, accuracy, sufficiency, non-infringement, or otherwise. All warranties, if any, are set forth in the applicable terms and conditions of sale accompanying the purchase of Mission Bio’s product. Without limiting the forgoing, Mission Bio and its affiliates provide no warranty and hereby disclaim any and all warranties as to the use of any third party products or protocols described herein. The use or resale of the products described herein may be subject to certain restrictions as set forth in the applicable terms and conditions of sale accompanying the purchase of such product. All products and services described herein are intended FOR RESEARCH USE ONLY and NOT FOR USE IN DIAGNOSTIC PROCEDURES.
QRC revised with AI assisted tools.