# Data Collector – DNA + RNA + Protein Workflow

Ensure Tapestri Instrument is updated with v3 firmware before use

| Chapter | Step | Question | Image/Data |
| --- | --- | --- | --- |
| 1 a ¦ Prepare Cell Suspension | A ¦ After first centrifugation step | Cell pellet visible? o Yes o No | [15 mL tube with cell pellet] |
|  | B ¦ After final washing step | Cell pellet visible? o Yes o No | [15 mL tube with cell pellet] |
|  | C ¦ First quantification | Cell concentration quantifiable? o Yes o No | [Cell suspension on slide, exported from imager] |
|  | D ¦ Quantification prior to cell staining | Cell concentration between 20,000 and 30,000 cells/µL in at least 40 µL? o Yes o No | [Cell suspension on slide, exported from imager, 5X dilution] |
| 1 b ¦ Stain and Wash Cells | E ¦ After first AOC washing step | Cell pellet visible? o Yes o No | [15 mL tube with cell pellet] |
|  | F ¦ After second AOC washing step | Cell pellet visible? o Yes o No | [15 mL tube with cell pellet] |
|  | G ¦ After third AOC washing step | Cell pellet visible? o Yes o No | [15 mL tube with cell pellet] |
|  | H ¦ Final quantification after cell dilution | Cell concentration between 4,000 and 4,500 cells/µL? o Yes o No | [Cell suspension on slide, exported from imager] |
| 2 ¦ Encapsulate Cells | A ¦ RNA mix preparation | RNA Lysis Mix diluted? o Yes o No | [Tube containing diluted RNA Lysis Mix with 20µL total volume] |
|  | B ¦ After reagent loading | Bubbles present in any of the reagents? o Yes o No | [Top view of DNA cartridge w/o gasket applied] |
|  | C ¦ When encapsulation program started | Encapsulation program started within 1 minute after loading RNA Mix + RNA lysis Mix o Yes o No | [Time between addition of lysis Mix and beginning of encapsulation program] |
|  | D ¦ After encapsulation program finished | Emulsions fully intact? o Yes o No | [Emulsion-safe tube w/ encapsulated cells] |
|  | E ¦ After encapsulation program finished | Remaining cell suspension volume [reservoir 2] < 10 µL? o Yes o No | [Measure volume] |
|  | F ¦ When removing oil | Total volume ~100 µL in the tube? o Yes o No | [Emulsion-safe tube w/ encapsulated cells and oil removed] |
| 3 ¦ Lysis, Protease Digest and RT | A ¦ Before thermal cycling | Thermal cycling protocol correct? o Yes o No | [Thermal cycling protocol image] |
|  | B ¦ Before thermal cycling | PCR skirt used? o Yes o No | [Thermal heat block image] |
|  | C ¦ After removing oil | Incubation started within 1 minute of completion of the encapsulation program o Yes o No | [Time between end of encapsulation and beginning of lysis program] |
|  | D ¦ After thermal cycling | Emulsions fully intact? o Yes o No | [Emulsion-safe tube w/ encapsulated cells] |
| 4 ¦ Barcode Cells | A ¦ After reagent loading, priming | Bubbles present in any of the reagents? o Yes o No | [Top view of DNA cartridge w/o gasket applied] |
|  | B ¦ After reagent loading, barcoding | Bubbles present in any of the reagents? o Yes o No | [Top view of DNA cartridge w/o gasket applied] |
|  | C ¦ After barcoding program finished | Emulsions fully intact and evenly distributed? o Yes o No | [8x emulsion-safe tubes w/ barcoded cells] |
|  | D ¦ After barcoding program finished | Record remaining Barcoding Bead volume [reservoir 7] o Yes o No | [Measure volume] |
|  | E ¦ After barcoding program finished | Record remaining Barcode Mix volume [reservoir 8] o Yes o No | [Measure volume] |
|  | F ¦ When removing oil | Total volume ~100 µL in each tube? o Yes o No | [8x emulsion-safe tubes w/ barcoded cells and oil removed] |
| 5 ¦ Targeted PCR Amplification | A ¦ Before thermal cycling | Thermal cycling protocol correct? o Yes o No | [Thermal cycling protocol without tubes, started with remaining time visible] |
|  | B ¦ Before thermal cycling | PCR skirt used? o Yes o No | [Thermal heat block image] |
|  | C ¦ After thermal cycling | Emulsions fully intact? o Yes o No | [8x emulsion-safe tubes w/ barcoded cells] |
| 6 ¦ Emulsion Breakage | A ¦ After emulsion breakage | Emulsions fully broken? o Yes o No | [8x emulsion-safe tubes w/ broken emulsions] |
| 7 ¦ Cleanup PCR Products | A ¦ After enzymatic cleanup | Sample spun down and transferred to new tube after enzymatic cleanup? o Yes o No | [Sample after transfer to new tube] |
|  | B ¦ After library quantification | Library concentration > 7 ng/µL o Yes o No | [Qubit reading] |
| 8 ¦ E-gel library extraction | A ¦ After electrophoresis | Product visible between 250-450 bp o Yes o No | [Gel before excision] |
|  | B ¦ After electrophoresis | Gel band excised between 250-450 bp o Yes o No | [Gel after excision] |
| 9 ¦ PCR Target Library | A ¦ After quantification w/ Qubit (Final Libraries) | Library concentration within spec? DNA: > 1.0 ng/µL o Yes o No RNA: > 1.0 ng/µL o Yes o No Protein: > 1.0 ng/µL o Yes o No | [Qubit reading] |
|  | B ¦ Bioanalyzer or TapeStation Results | Expected peak at ~460 bp (DNA/RNA libraries) and ~250 bp (protein library) with few (<5%) off-target fragments (e.g., primer dimers)? o Yes o No | [Fragment analyzer results] |