Rev A

QUICK REFERENCE CARD

Targeted Single-cell DNA + RNA Fusion + Protein Sequencing Using the Tapestri® Platform v3

# IMPORTANT

- Ensure instrument is using the current v3 firmware.

- Vortex all reagents unless directed otherwise.

- Thaw -20 °C reagents on ice.

- Avoid sources of static and pipette slowly and carefully when handling emulsions.

- Provide debris-free cell suspension with >90% viability.

- Starting cell concentration needs to be 25,000 cells/μL..

- Final cell concentration needs to be ~3,000 cells/μL

- Equilibrate Encapsulation and Barcoding oils for 30 minutes prior to use.

- Always use a PCR skirt.

# Stain Cells With Antibody Panel

## RECONSTITUTE LYOPHILIZED ANTIBODY PANEL

- Equilibrate BioLegend lyophilized Heme panel at room temperature for 5 minutes.

- Spin down at 10,000 x g for 30 seconds at room temperature.

- Resuspend lyophilized panel in 60 μL of Cell Staining Buffer (CSB). Vortex for 10 seconds.

- Incubate for 5 minutes at room temperature.

- Vortex for 10 seconds and spin down at 10,000 x g for 30 seconds at room temperature.

- Transfer entire 60 μL to a 1.5 mL Protein LoBind Eppendorf tube.

- Centrifuge at 14,000 x g for 15 minutes at 4°C. In the meantime, proceed to the next step.

## BLOCK CELLS

- Using CSB, dilute cells to 25,000 cells/μL in a total volume of 40 μL.

- In a 15 mL DNA LoBind conical tube, gently add the following reagents:

| Reagent | Volume (μL) |
| --- | --- |
| Cell Suspension (25,000 cells/μL) | 40 |
| Blocking Buffer | 5 |
| Human TruStain FcXTM | 5 |
| Total Volume | 50 |

- Mix by pipetting up and down with a wide-bore P-200 pipette and incubate on ice for 15 minutes.

## STAIN CELLS

- Using a P-200 pipette, add 50 μL of the resuspended panel (Step 7) to the blocked cell suspension for a total volume of 100 μL.

IMPORTANT   Avoid touching the bottom or sides of the tube with the pipette tip to avoid pelleted protein aggregates.

- Mix by pipetting up and down and incubate on ice for 30 minutes.

## WASH CELLS

IMPORTANT   Equilibrate Encapsulation Oil for 30 minutes prior to use.

- Add 14 mL of pre-chilled CSB to the cell staining solution.

- Centrifuge at 400 x g for 10 minutes at 4 °C in a swinging bucket.

- Carefully aspirate 13.5 mL of supernatant using a serological pipette. Leave 500 μL of supernatant behind. Do not disturb the pellet.

- Repeat Steps 13 - 15 for two additional washes (centrifuge for 5 minutes instead of 10 minutes).

- Remove 13 mL of supernatant using a serological pipette.

- Remove 400 μL of supernatant using a P-1000 pipette, leaving 100 μL.

- Add 900 μL of DPBS to the cell pellet. Gently resuspend by pipetting up and down with a wide-bore P-1000 pipette.

- OPTIONAL: Filter the cells with a 40 μM Flowmi cell strainer and transfer to a 1.5 mL DNA LoBind Eppendorf tube.

- Centrifuge at 400 x g for 5 minutes at 4 °C.

- Carefully remove all supernatant. Do not disturb the pellet.

- Resuspend the pellet in 60 μL of DPBS by pipetting up and down and count cells.

- If necessary, use DPBS to dilute cells to 2,800 - 3,200 cells/µL.

Cell concentrations as low as 1,000 cells/µL may be acceptable.

# Encapsulate Cells

- Mount the Base Plate onto the Tapestri Instrument.

- Place 0.2 mL emulsion-safe PCR tube into the middle of the slot at the left side of the Base Plate.

- Place the DNA cartridge onto the base plate.

- Preheat the thermal cycler with the Lysis/Digest protocol (hold at 50C).

| Step | Temperature | Time |
| --- | --- | --- |
| 1 | 50 °C | 120 min |
| 2 | 80 °C | 10 min |
| 3 | 4 °C | HOLD |

- Preheat the RNA Fusion rev primer pool to 70C for 10 minutes, then store on ice until ready to use.

- In a 0.2 mL PCR tube, combine 2 μL RNA Lysis Mix with 18 μL nuclease-free water. Gently mix by pipetting up and down 5 times. Do not vortex. Store on ice.

- In a new tube, prepare RT + Lysis Mix:

| Reagent | Volume (μL) |
| --- | --- |
| RNA Enzyme Buffer | 12 |
| RNA Reagent 1 | 3 |
| RNA Reagent 3 | 3 |
| RNA Reagent 4 | 3 |
| RNA Reagent 2 | 5.4 |
| RNA Enzyme | 3 |
| RNA Fusion reverse primer pool | 6 |
| Nuclease-free water | 23.3 |
| Diluted RNA Lysis Mix | 1.3 |
| Total Volume | 60 |

IMPORTANT   Add the Diluted RNA Lysis Mix immediately before loading the cartridge.

- Pipette 60 μL of Lysis Mix into

reservoir 1.

- Gently pipette mix the Cell Suspension

then pipette 35 uL (3,000 cells/μL) into reservoir 2.

- Pipette 200 μL of Encapsulation Oil into reservoir 3.

- Apply DNA Gasket on top of the cartridge and close the instrument lid.

- Run the Cell Encapsulation program by pressing Step 1: Encapsulation on the Tapestri Instrument touch screen.

- Once program completes, press DONE and remove the cartridge and collection tube from the Base Plate.

- Use a gel loading tip to carefully remove excess oil from the bottom layer of the tube leaving a total of 100 µL of emulsion + oil

(approximately the middle point of the tube as it narrows to the conical base).

IMPORTANT   Do not stop here. Immediately move on to RT + Lysis.

# RT + Lysis

- Immediately place the sample tube into the preheated thermocycler and run the Lysis/Digest protocol.

| Step | Temperature | Time |
| --- | --- | --- |
| 1 | 50 °C | 120 min |
| 2 | 80 °C | 10 min |
| 3 | 4 °C | HOLD |

- Once complete, store the lysed and digested sample at 4 °C until required in the next step. Proceed within 2 hours.

# Barcode Cells

PRIME THE DNA CARTRIDGE

- Place eight emulsion-safe PCR tubes

into the eight slots at the bottom of the Base Plate.

- Mount the DNA Cartridge back onto the Base Plate.

- Pipette 200 μL of Electrode Solution into each reservoir 4 and 500 μL of Electrode Solution into each reservoir 5.

- Apply DNA Gasket on top of the cartridge and close instrument lid.

- Run the Priming program by pressing Step 2: Barcoding on the Tapestri Instrument touch screen.

- In the meantime, retrieve the Barcoding Oil and Barcoding Beads v3 (•) and leave at room temperature (protect from light) and prepare 300 μL Barcode Mix:

| Reagent | Volume (μL) |
| --- | --- |
| Barcoding MM v3 | 289.9 |
| DNA forward primer pool | 2.5 |
| DNA reverse primer pool | 3.1 |
| RNA Fusion forward primer pool | 2.5 |
| Antibody Tag Primer (l) | 2 |
| Total Volume | 300 |

- Briefly vortex the Barcode Mix, quick-spin to collect the contents, and store on ice.

LOAD THE DNA CARTRIDGE

- Quick-spin Barcoding Beads v3 to collect contents. Take 67 μL of the prepared Barcode Mix and add it to the barcoding bead tube.

- Retrieve the lysed and digested sample from the thermocycler.

- Slowly pipette the entire cell encapsulation sample (~100 μL), including any oil at the bottom into reservoir 6.

- Vortex Barcoding Beads v3 at full speed for 1 minute. Pipette 250 μL of Barcoding Beads v3 into reservoir 7.

- Pipette 200 μL of Barcode Mix into reservoir 8.

- Pipette 1.25 mL of Barcoding Oil into reservoir 9.

- Apply DNA Gasket on top of the cartridge and close the instrument lid.

- Run the Cell Barcoding program by pressing NEXT on the Tapestri Instrument touch screen.

- Once the program completes, press DONE and remove the Base Plate together with the cartridge and the eight collection tubes.

- Remove the DNA Cartridge from the Base Plate.

- Use a gel loading tip to carefully remove excess oil from the bottom layer of all eight tubes, leaving a total of 100 µL of emulsion + oil (approximately the middle point of the tube as it narrows to the conical base) per tube.

# Targeted PCR Amplification

- Transfer the sample tubes to the thermocycler, and run the Targeted PCR protocol based on the table located on the next page

| Amplicon Number |  |  | 20-100 | 101-200 | 201-300 | >300 |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Step | Ramp | Temp | Time | Time | Time | Time | Cycle |
| 1 | 4 °C/s | 98 °C | 6 min | 6 min | 6 min | 6 min |  |
| 2 | 1 °C/s | 95 °C | 30 sec | 30 sec | 30 sec | 30 sec | 11 |
| 3 |  | 72 °C | 10 sec | 10 sec | 10 sec | 10 sec |  |
| 4 |  | 61 °C | 3 min | 4.5 min | 6 min | 9 min |  |
| 5 |  | 72 °C | 20 sec | 20 sec | 20 sec | 20 sec |  |
| 6 | 1 °C/s | 95 °C | 30 sec | 30 sec | 30 sec | 30 sec | 13 |
| 7 |  | 72 °C | 10 sec | 10 sec | 10 sec | 10 sec |  |
| 8 |  | 51 °C | 3 min | 4.5 min |  |  |  |
| 9 |  | 72 °C | 20 sec | 20 sec | 20 sec | 20 sec |  |
| 10 | 4 °C/s | 72 °C | 2 min | 2 min | 2 min | 2 min |  |
| 11 |  | 4 °C | HOLD | HOLD | HOLD | HOLD |  |

IMPORTANT   Make sure to set up two separate cycle programs (11 cycles and 13 cycles)  each with a ramp rate of 1 ºC/s.

BREAK EMULSIONS

- Add 15 μL of Extraction Agent  to each tube. Briefly vortex and spin for 20 seconds.

- Incubate at room temperature for 3 minutes.

- Add 50 μL of nuclease-free water to each tube.

- Briefly vortex and spin for 10 seconds.

- Pipette the total volume of the aqueous top layer from each tube into a new 1.5 mL DNA LoBind Eppendorf tube. Do not transfer any oil or Barcoding Beads.

- Spin down 1.5 mL tubes at 3000 g for 5 minutes. There should be two layers, the clear aqueous layer on top and an opaque layer on the bottom.

- Carefully transfer the clear supernatant from each tube into a new 1.5 mL DNA LoBind Eppendorf tube.

- Bring the total volume in each tube up to 336 uL with nuclease-free water.

- Proceed to next step or store at -20 °C.

# Digest PCR Product

- Add 40 uL DNA Clean Up Buffer and 24 uL DNA Clean Up Enzyme to the sample tube for a total of 400 uL per tube.

- Briefly vortex and spin down the tube.

- Transfer tube to a thermo mixer or heat block and incubate at 37 °C for 60 minutes.

- Immediately continue to the next step.

# Clean Up PCR Product

NOTE   Equilibrate AMPure XP reagent to room temperature.

Prepare 5 mL fresh 80% ethanol using nuclease-free

water.

### Spin down sample tube for 20 seconds. If a pellet is visible, transfer clear aqueous solution to a new tube, being careful not to disturb the pellet. Add nuclease-free water if any evaporation occurred for a total volume of 400 μL.

- Thoroughly vortex AMPure XP reagent at high speed immediately prior to usage.

- Add 280 μL (0.70X) of AMPure XP reagent to tube. Vortex for 5 seconds and quick-spin to collect contents.

- Incubate tube at room temperature for 5 minutes.

- Place on magnet and wait 5 minutes for the beads to separate from solution.

IMPORTANT  Do not discard the supernatant from the tube as it contains the protein library.

- Without removing the tube from the magnet, transfer the supernatant (~680 μL) from tube to a new 1.5 mL DNA LoBind Eppendorf tube and set aside at room temperature for Protein Amplicon Cleanup.

- Carefully add 800 uL of the freshly prepared 80% ethanol, wait 30 seconds. Without removing the tube from the magnet, remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 77 once, for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from tube without disturbing the beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 5 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet. Add 50 uL of low TE into tube. Vortex and quick-spin to collect the contents.

- Incubate the tube at room temperature for 2 minutes.

- Place the tube onto the magnet and wait for at least 2 minutes or until the solutions are clear.

- Transfer 50 μL of purified PCR product to a new 0.2 mL PCR tube.

# Separate DNA and RNA Fusion amplicons

## RNA Fusion AMPLICON CLEANUP

PREPARE STREPTAVIDIN BEADS

- Thoroughly vortex Streptavidin beads. Transfer 60 μL of Streptavidin Beads to a new 1.5 mL DNA LoBind Eppendorf tube.

- Place on a magnet and wait 2 minutes for the beads to separate from the solution.

- Remove the supernatant and discard.

- Wash the beads with 1 mL of nuclease-free water. Wait 1 minute with the tube on the magnet for solution to clear then remove the supernatant.

- Repeat Step 88 once for a total of two washes.

- Resuspend the beads in 30 μL of 1X Wash Buffer (listed below).

| Reagent | Final concentration |
| --- | --- |
| Tris-HCl pH 7.5 | 5 mM |
| EDTA | 0.5 mM |
| NaCl | 1 M |

ISOLATE RNA Fusion AMPLICONS

- Retrieve the tube from the Clean Up PCR Product section that contains the DNA and RNA Fusion amplicons.

- Add 4 μL of TS Biotin oligo (20 uM Biosg-TS). Briefly vortex and quick-spin.

- Incubate at 96 °C for 5 minutes.

- Transfer the tubes immediately onto ice and incubate for 5 minutes.

- Add and mix 30 μL of Streptavidin Beads resuspended in 1X Wash Buffer to TS Biotin oligo-treated sample tube.

- Incubate for 30 minutes at room temperature.

- Place on magnet, wait 5 minutes for the beads to separate from the solution.

- Transfer the supernatant to a new 0.2 mL PCR tube. This tube contains the DNA amplicons.

- Wash the Streptavidin Beads with 16 uL of nuclease-free water and add this supernatant to the 0.2 mL PCR tube containing the DNA amplicons.

- Resuspend the streptavidin beads in 15 μL of nuclease-free water. The RNA Fusion amplicons are bound to the streptavidin beads. Proceed to Library PCR or store at 4 °C.

DNA AMPLICON CLEANUP

- Take the 0.2 mL PCR tube containing the DNA amplicons and bring up to 100 uL with nuclease-free water if needed.

- Add 72 μL (0.72X) of AMPure XP reagent to the tube with the DNA amplicons.

- Vortex for 5 seconds and quick-spin to collect the contents.

- Incubate tube at room temperature for 5 minutes.

- Place on magnet, wait 5 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of the freshly prepared 80% ethanol, wait 30 seconds. Without removing the tube from the magnet, remove the ethanol without disturbing the AMPure XP beads

- Repeat Step 107 once, for a total of two washes

- Keeping the tube on the magnet, remove all residual ethanol from tube without disturbing the beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet. Add 15 μL of nuclease-free water. Vortex and quick-spin to collect the contents.

- Incubate the tube at room temperature for 2 minutes.

- Place the tube onto the magnet and wait for at least 2 minutes or until the solutions are clear.

- Transfer 15 μL of purified DNA amplicons to new 0.2 mL PCR tube.

- Store the purified PCR product on ice and proceed to the next step, or store at -20 °C long term.

## PROTEIN AMPLICON CLEANUP

PREPARE STREPTAVIDIN BEADS

- Thoroughly vortex Streptavidin beads. Transfer 100 μL of Streptavidin Beads to a new 1.5 mL DNA LoBind Eppendorf tube.

- Place on a magnet and wait 2 minutes for the beads to separate from the solution.

- Remove the supernatant and discard. Wash the beads with 1 mL of 2X Wash Buffer. Wait 1 minute for solution to clear.

- Repeat Step 118 once for a total of two washes.

- Remove the supernatant and resuspend the beads in 690 μL of 2X Wash Buffer. Set aside until Step 125 below.

ISOLATE ANTIBODY TAGS

- Retrieve the tube with the supernatant from the Clean Up PCR Product section that contains the protein library. Split the solution into two 1.5 mL DNA LoBind Eppendorf tubes (340 μL each).

- To each tube, add 2 μL of Biotin Oligo (●) to the supernatant. Briefly vortex and quick-spin.

- Incubate at 96 °C for 5 minutes.

- Transfer the tubes immediately onto ice and incubate for 5 minutes.

- Add and mix 342 μL of Streptavidin Beads resuspended in 2X Wash Buffer - from Step 120 above - to each Biotin Oligo-treated sample tube.

- Incubate for 20 minutes on a shaker at room temperature.

- Place on magnet, wait 5 minutes for the beads to separate from the solution.

- Remove the supernatant and wash the Streptavidin Beads with 1 mL of 1X Wash Buffer.

- Discard the supernatant, remove tubes from the magnet and wash a second time with 1 mL nuclease-free water. Mix by pipetting up and down 5 times.

- Place on the magnet, wait 3 minutes for beads to separate from the solution.

- Remove the supernatant. In each tube, resuspend the beads in 25 μL of nuclease-free water. Transfer and combine into a new 0.2 mL PCR tube (50 μL). The Protein library is bound to the Streptavidin Beads (brown). Proceed to Library PCR or store at 4 °C.

# Library PCR

- Set up three different Library PCR reactions, one for the DNA Library, one for the RNA Fusion Library, and one for the Protein Library as follows:

IMPORTANT  	Ensure DNA Library Index Primers are used for DNA, TS Library Index Primers are used for RNA Fusion, and Protein Library Index Primers are used for protein..

|  | DNA | RNA Fusion | Protein |
| --- | --- | --- | --- |
| Reagent | Volume [μL] |  |  |
| Library Mix | 25 | 25 | 25 |
| DNA Library Index Primer (●) | 10 | - | - |
| TS RNA Fusion Library i7 Primer | - | 5 | - |
| Library i5 Primer | - | 5 | - |
| Protein Library Index Primer (●) | - |  | 10 |
| DNA amplicons | 15 | - | - |
| Resuspended Streptavidin Beads containing RNA Fusion amplicons | - | 15 | - |
| Resuspended Streptavidin Beads containing Antibody Tags | - | - | 15 |
| Total Volume | 50 | 50 | 50 |

- Vortex and quick-spin to collect the contents.

- Transfer the sample tubes to a thermocycler and run the following Library PCR protocol:

| Step | Temperature | Time | Cycle |
| --- | --- | --- | --- |
| 1 | 95 °C | 3 min |  |
| 2 | 98 °C | 20 sec | 12 for DNA Library 14 for RNA Fusion Library 20 for Protein Library |
| 3 | 62 °C | 20 sec |  |
| 4 | 72 °C | 45 sec |  |
| 5 | 72 °C | 2 min |  |
| 6 | 4 °C | HOLD |  |

# Clean Up Library PCR Product

NOTE  	Equilibrate AMPure XP reagent to room temperature. Thoroughly vortex AMPure XP reagent for 15 seconds at high speed.

DNA LIBRARY CLEANUP

- Add nuclease-free water to the sample tubes containing the DNA library to bring up to 50 uL if evaporation occurred.

- Add 0.69X of AMPure XP reagent to the 50 μL DNA library tubes. Vortex for 10 seconds and quick-spin to collect the contents.

- Incubate the tube at room temperature for 5 minutes.

- Place on the magnet, wait 5 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of freshly prepared 80% ethanol, wait 30 seconds, and remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 139 once for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from the tube without disturbing the AMPure XP beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet and add 50 μL of nuclease-free water. Vortex and quick-spin to collect the contents.

- Incubate for 2 minutes.

- Place the tubes on the magnet and wait 2 minutes or until the solution is clear.

- Transfer 50 μL of the DNA library to a new 0.2 mL PCR tube.

- Add 0.69X of AMPure XP reagent to the 50 μL sample tube.

- Incubate the tube at room temperature for 5 minutes.

- Place on the magnet, wait 5 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of freshly prepared 80% ethanol, wait 30 seconds, and remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 151 once for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from the tube without disturbing the AMPure XP beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet and add 12 μL low TE. Vortex and quick-spin to collect the contents.

- Incubate the tube at room temperature for 2 minutes.

- Place the tube onto the magnet and wait for at least 2 minutes or until the solution is clear.

- Transfer 12 μL of DNA library to a new 0.2 mL PCR tube or 1.5 mL DNA LoBind Eppendorf tube. Store at -20 °C for long-term storage.

RNA Fusion LIBRARY CLEANUP

- Place the tubes containing RNA Fusion Library on the magnet and wait 2 minutes for Streptavidin Beads to separate from the solution.

- Without removing the tube from the magnet, transfer 50 μL of supernatant to  new 0.2 mL PCR tubes.

- Add 0.69X of AMPure XP reagent to the 50 μL sample tube.

Vortex for 10 seconds and quick-spin to collect the contents.

- Incubate the tube at room temperature for 5 minutes.

- Place on the magnet and wait 2 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of freshly prepared 80% ethanol, wait 30 seconds, and remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 166 once for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from the tube without disturbing the AMPure XP beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet and add 50 μL of nuclease-free water. Vortex and quick-spin to collect the contents.

- Incubate for 2 minutes.

- Place the tubes on the magnet and wait 2 minutes or until the solution is clear.

- Transfer 50 μL of the RNA Fusion library to a new 0.2 mL PCR tube.

- Add 0.69X of AMPure XP reagent to the 50 μL sample tube.

- Incubate the tube at room temperature for 5 minutes.

- Place on the magnet, wait 2 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of freshly prepared 80% ethanol, wait 30 seconds, and remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 178 once for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from the tube without disturbing the AMPure XP beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet and add 12 μL low TE. Vortex and quick-spin to collect the contents.

- Incubate the tube at room temperature for 2 minutes.

- Place the tube onto the magnet and wait for at least 2 minutes or until the solution is clear.

- Transfer 12 μL of RNA Fusion library to a new 0.2 mL PCR tube or 1.5 mL DNA LoBind Eppendorf tube. Store at -20 °C for long-term storage.

## PROTEIN LIBRARY CLEANUP

- Place the tube (Protein Library) on the magnet and wait 2 minutes for Streptavidin Beads to separate from the solution.

- Without removing the tube from the magnet, transfer 50 μL of supernatant to a new 0.2 mL PCR tube.

- Add 45 μL (0.90X) of AMPure XP reagent to the 50 μL sample tube.

Vortex for 10 seconds and quick-spin to collect the contents.

- Incubate the tube at room temperature for 5 minutes.

- Place on the magnet and wait 2 minutes for the beads to separate from the solution.

- Without removing the tube from the magnet, remove the supernatant and discard.

- Carefully add 180 μL of freshly prepared 80% ethanol, wait 30 seconds, and remove the ethanol without disturbing the AMPure XP beads.

- Repeat Step 192 once for a total of two washes.

- Keeping the tube on the magnet, remove all residual ethanol from the tube without disturbing the AMPure XP beads.

- Dry AMPure bead pellets in the tube on the magnet by incubating at room temperature for 2 minutes. Avoid overdrying the beads.

- Remove the tube from the magnet and add 17 μL of nuclease-free water. Vortex and quick-spin to collect the contents.

- Incubate the tube at room temperature for 2 minutes.

- Place the tube onto the magnet and wait for 2 minutes or until the solution is clear.

- Transfer 15 μL of purified PCR product from the tube to a new 0.2 mL PCR tube or 1.5 mL DNA LoBind Eppendorf tube. Store at -20 °C for long-term storage.

# Quantify and Pool Library

- Run 1 μL of each library on a High-Sensitivity Bioanalyzer chip or equivalent. Repeat AMPure cleanups if necessary. The libraries can be reamplified if necessary for enough material for additional cleanups. Once the libraries are clean, quantify by qPCR. Normalize libraries to 2 nM and pool using the Tapestri Sample Pooling tool PN40676.

Contact support@missionbio.com if final libraries are < 2.0 ng/μL

IMPORTANT 	Refer to the Tapestri Single-Cell DNA

Sequencing v3 User Guide (MB05-0017) for

additional information on how to accurately quantify Tapestri libraries that include large size off-target products.

# Sequence Tapestri Single-cell DNA and RNA Fusion Libraries

## IMPORTANT

- The final DNA library consists of target-specific amplicons ranging from 350 - 550 bp with a peak at 460 bp.

- The final RNA Fusion library consists of target-specific amplicons ranging from 350 - 700 bp with a peak at 460 bp.

- The final Protein library consists of target-specific amplicons ranging from 230 - 270 bp with a peak at 250 bp.

- Refer to the Tapestri Single-Cell DNA Sequencing v3 User Guide (MB05-0017) for sequencing recommendations.

For technical support visit www.missionbio.com/support or email support@missionbio.com.

© 2023 Mission Bio, Inc. All rights reserved. Duplication, public display, public performance, modification, or any other use of all or any portion of this document without the express prior written consent of Mission Bio is strictly forbidden. “Mission Bio” and the “MB” logo are trademarks of Mission Bio and may not be used for any purpose without the express prior written consent of Mission Bio. All other trademarks are the property of their respective owners. Nothing in this documentation shall constitute a warranty, express or implied, as to the products (whether a product of Mission Bio, its affiliates, or any third party) or any protocols described herein, whether as to performance, accuracy, sufficiency, non-infringement, or otherwise. All warranties, if any, are set forth in the applicable terms and conditions of sale accompanying the purchase of Mission Bio’s product. Without limiting the forgoing, Mission Bio and its affiliates provide no warranty and hereby disclaim any and all warranties as to the use of any third party products or protocols described herein. The use or resale of the products described herein may be subject to certain restrictions as set forth in the applicable terms and conditions of sale accompanying the purchase of such product. All products and services described herein are intended FOR RESEARCH USE ONLY and NOT FOR USE IN DIAGNOSTIC PROCEDURES.
QRC revised with AI assisted tools.